## Slide 1
Introduction
Junsub Kim1, Jinsoo Park1, Muhwan Kim2, Yeonsook Jang2, Somasekhar R. Chinnadayyala1,3 and Sungbo Cho1,4*
1 Department of Health Sciences and Technology, GAHIST, Gachon University, Incheon, Korea
2 Femto Science Inc., #505, Kumkang Penterium IT Valley, 557, Dongtangiheung-Ro, Hwaseong, Gyeonggi-do, Korea
3 Department of Biomedical Engineering, Gachon University, Incheon, Korea
4 Department of Electronic Engineering, Gachon University, Seongnam, Korea
Atmospheric-Pressure and Low-Temperature Plasma Needle
for Biochip Applications
Atmospheric-pressure and low-temperature plasmas have been employed for diverse biomedical applications including the surface treatment of medical devices, inactivation of microorganisms, or the sterilization. Hydrophilic surface modification of the biochip is widely required to immobilize DNA or antibody probes for disease diagnosis, and to pattern the cells for high throughput screening of drug candidates.
In this study, we developed a plasma needle for atmospheric-pressure and low-temperature plasma treatment and evaluated it for hydrophilic surface modification. The plasma needle was constructed using a stainless syringe needle connected to a low frequency power supply for floating at high electrical potential with respect to the ground plate electrode. While the argon gas as working gas was injected through the needle, the plasma was generated and propagated to open air from the needle tip.
The plasma needle was applied to the inlet of the fabricated poly(dimethylsiloxane) microchannel which put on the hydrophobic polystyrene dish. The area of the polystyrene dish treated by atmospheric pressure and low temperature plasma was measured by measuring the contact angle of a water drop and by observing the adhesion and growth of HEK 293/GFP cells. As a result, it was confirmed that his area exhibited a hydrophilic state.
Results
Conclusion
The developed plasma needle is able to apply the atmospheric-pressure and low-temperature plasma for microfluidic biochip applications.
The hydrophobic surface of the polystyrene petri dish is changed to the hydrophilic state by treating with the plasma needle.
The HEK 293/GFP cells adhere and grow on the plasma-treated polystyrene dish by using the developed plasma needle.
The work was supported by the National Research Foundation of Korea (No. NRF-2017R1D1A1A09000712, 2018R1C1B6009385, 2018M3A9F1023690).
Acknowledgement
Developed Plasma Needle System
Methodology
Experimental Procedure
1. Coating of SU-8 (5 μm) on Si-wafer
2. UV exposure
(Power: 360mW, Wavelength: 365 nm)
3. Development of SU-8
5. Fabrication of PDMS Mold
4. Molding of PDMS
Microfluidic Chip
Microfluidic chip was fabricated to evaluate the hydrophilic surface modification of plasma needle.
Microchannel with a height of 5 μm, width of 40 μm
(a)
(b)
Fabricated SU-8 mold (a), micro-fluidic devise discharge test using bule and red ink (b), (c)
Phase contrast & Fluorescent microscope images
200 μm
200 μm
(c)
(d)
Phase contrast microscope images of HEK 293/GFP cell cultured 3 days on Ar plasma treated porystyrene petri dish (c), and fluorescent image (d)
8. Removing PDMS mold
x40
x40
Hydrophilicity versus Ar Plasma Treatment Time
ɵ
γSL
γS
γL
γS = γSL + γL cosθ
γS : Free energy of interface between solid / gas
γSL : Free energy of interface between liquid/ solid
γL : Free energy of interface between liquid / gas
ɵ = 0 ° : Liquid spreads on the solid surface completely
0° < ɵ < 90 ° : Liquid wetted on solid surface
90° < ɵ < 180 ° : Liquid don’t wetted on solid surface
ɵ = 180 ° : Liquid don’t wetted on solid surface completely
0sec
30sec
90sec
180sec
Phase contrast microscope images of HEK 293/GFP cell cultured 3 days on hydrophobic polystyrene petri dish (a), and fluorescent image (b)
(a)
(b)
200 μm
x40
200 μm
x40
6. Putting PDMS mold on porystyrene petri dish
7. Ar Plasma treatment
Ar Plasma
Ground
9. Culturing HEK 293/GFP cell
Ar
DC Power supply
Transformer
Main board
Ground
Ar Plasma
Plasma needle
Ar Plasma treatment
Pressure : 0.3MPa (2.96Torr)
Power : 8w
Time : 40s
+8v
x40
200 μm
(c)