## 하네스
Battery
Pin-map | [5 empty] | Part List
| Pin No | Function |  | Pin No | Function | PCB | PN | Vendor | Description
Main PCB | 1 | GND | Battery | 1 | GND | Main | 053048-0310 | Molex | 1.25mm Pitch DIP CON, Right Angle 3-Pin
| 2 | Battery Temp ADC |  | 2 | Battery Temp ADC | 하네스 | 051021-0300 | Molex | 1.25mm Pitch Housing, Female, 3-Pin
| 3 | Battery Vcc |  | 3 | Battery Vcc |  | 050058-8000 | Molex | 1.25mm Pitch Crimp Terminal, Female, 28-32 AWG
[7 empty] | 케이블 |  | AWG28
Battery
Transformer
Pin-map | [5 empty] | Part List
| Pin No | Function |  | Pin No | Function | PCB | PN | Vendor | Description
Main PCB | 1 | Gas S/V Enable | S/V | 1 | Gas S/V Enable | Main | 053048-0710 | Molex | 1.25mm Pitch DIP CON, Right Angle 7-Pin
| 2 | GND |  | 2 | GND | 하네스 | 051021-0700 | Molex | 1.25mm Pitch Housing, Female, 7-Pin
| 3 | Plasma On Key | Transformer PCB | 1 | Plasma On Key |  | 050058-8000 | Molex | 1.25mm Pitch Crimp Terminal, Female, 28-32 AWG
| 4 | GND |  | 2 | GND |  | 케이블 |  | AWG28
| 5 | PULSE_OUT2 |  | 3 | PULSE_OUT2 |  | 050058-8000 | Molex | 1.25mm Pitch Crimp Terminal, Female, 28-32 AWG
| 6 | PULSE_OUT1 |  | 4 | PULSE_OUT1 |  | 051021-0500 | Molex | 1.25mm Pitch Housing, Female, 5-Pin
| 7 | +8.0V |  | 5 | +8.0V | Transformer | 053398-0571 | Molex | 1.25mm Pitch SMD CON, Vertical 5-Pin
GND 하네스
Pin-map | [5 empty] | Part List
| Pin No | Function |  | Pin No | Function | PCB | PN | Vendor | Description
Main PCB | 1 | PAD - GND |  |  |  | Main | DC-011C_SMD |  | DC Power Jack, 0.7 Pie, DC Plug, SMD
| 2 | DC Jack detect Pin |  |  |  | 하네스 | NA645 | 라이트컴 | Coms 컨넥터-DC파워 2.5Ø x 0.7(각형)
| 3 | Case Pin - GND | GND Cable | 1 | Case Pin - GND |  | 케이블 |  | AWG20
| 4 | Center Pin - GND |  | 2 | Center Pin - GND |  | 악어클립 |  | 악어클립 대(검정색)
| 5 | PAD - GND
| 6 | Hole
| 7 | Hole
JTAG 하네스
Pin-map
| Pin No | Function |  | Pin No | Function
Main PCB | 1 | USB VBUS | ST-LINK | 1,2 | MCU VDD
| 2 | SYS_SWDIO |  | 7 | SYS_SWDIO
| 3 | SYS_SWCLK |  | 9 | SYS_SWCLK
| 4 | USB_ID |  | 3 | GND
| 5 | GND |  | 4 | GND

## ME Size
Battery 용량
| 용량 | W | L | H | 체적
| 750 | 48 | 45 | 3 | 6480
|  | 20 | 90 | 10 | 18000 | 2.7778 | 2083.3333
|  | W | L | H
| PCB | 24 | 68 | 1
| Battery | 20 | 90 | 10
ME | Main | 30 | 100 | 30
| Sub | 30 | 100 | 30
부품 높이 제한
|  | Top | 4 | mm
|  | Bottom | 9 | mm

## Battery CON
|  | Body | Spring | Total
9155-000 | AVX | 3 | 2.3 | 5.3
K113868632 | Top-Link |  |  | 5.3
| Tyco |  |  | 6.2
|  | 3.45 | 1.45 | 4.9
0.089
[5 empty] | 0.204 | 0.182
[5 empty] | 0.115 | 0.11
0.072

## Power review
각 Stage별 전압/전류/파워 | [5 empty] | Battery 사용 시간
Battery | V_batt | 4 | 4 | 4 | V |  | Battery 용량 | 1000 | mAh
| I_Peak | 2 | 2.5 | 2.5 | A |  | I_Peak | 2 | 2 | 2 | 2 | 2 | A
| Power_in | 8 | 10 | 10 | W |  | On Time | 20 | 40 | 60 | 80 | 100 | %
Booster | V_in | 8 | 8 | 8 | V |  | I_average | 400 | 800 | 1200 | 1600 | 2000 | mA
| I_in | 1 | 1.25 | 1.25 | A |  | 사용시간 | 2.5 | 1.25 | 0.83333 | 0.625 | 0.5 | Hour
Transformer | Turn ratio | 125 | 392 | 350
| V_out | 1000 | 3136 | 2800 | V |  | Battery 용량 | 2000 | mAh
| I_out | 0.008 | 0.0031888 | 0.0035714 | W |  | I_Peak | 2 | 2 | 2 | 2 | 2 | A
[7 empty] | On Time | 20 | 40 | 60 | 80 | 100 | %
[7 empty] | I_average | 400 | 800 | 1200 | 1600 | 2000 | mA
[7 empty] | 사용시간 | 5 | 2.5 | 1.6667 | 1.25 | 1 | Hour
[7 empty] | Battery 용량 | 2000 | mAh
[7 empty] | I_Peak | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | A
| 280 | 250 |  |  |  |  | On Time | 20 | 40 | 60 | 80 | 100 | %
| 7 | 7 |  |  |  |  | I_average | 500 | 1000 | 1500 | 2000 | 2500 | mA
| 1960 | 1750 |  |  |  |  | 사용시간 | 4 | 2 | 1.3333 | 1 | 0.8 | Hour
| 392 | 350

## LED
| V1.0 | V2.0
VCC | 3.3 | 3.3 | V
R | 470 | 180
I | 7.0213 | 18.3333 | mA