## 하네스
Battery
Pin-map | [5 empty] | Part List
| Pin No | Function |  | Pin No | Function | PCB | PN | Vendor | Description
Main PCB V1.0 | 1 | GND | Battery | 1 | GND | Main | 053048-0310 | Molex | 1.25mm Pitch DIP CON, Right Angle 3-Pin
| 2 | Battery Temp ADC |  | 2 | Battery Temp ADC | 하네스 | 051021-0300 | Molex | 1.25mm Pitch Housing, Female, 3-Pin
| 3 | Battery Vcc |  | 3 | Battery Vcc |  | 050058-8000 | Molex | 1.25mm Pitch Crimp Terminal, Female, 28-32 AWG
[7 empty] | 케이블 |  | AWG28
| Pin No | Function |  | Pin No | Function |  | Battery
Main PCB V2.0 | 3 | GND | Battery | 1 | GND
| 2 | Battery Temp ADC |  | 2 | Battery Temp ADC | Part List
| 1 | Battery Vcc |  | 3 | Battery Vcc | PCB | PN | Vendor | Description
[6 empty] | Main | 053048-0710 | Molex | 1.25mm Pitch DIP CON, Right Angle 7-Pin
[6 empty] | 하네스 | 051021-0700 | Molex | 1.25mm Pitch Housing, Female, 7-Pin
Transformer | [6 empty] | 050058-8000 | Molex | 1.25mm Pitch Crimp Terminal, Female, 28-32 AWG
Pin-map | [6 empty] | 케이블 |  | AWG28
| Pin No | Function |  | Pin No | Function |  | 050058-8000 | Molex | 1.25mm Pitch Crimp Terminal, Female, 28-32 AWG
Main PCB V1.0 | 1 | Gas S/V Enable | S/V | 1 | Gas S/V Enable |  | 051021-0500 | Molex | 1.25mm Pitch Housing, Female, 5-Pin
| 2 | GND |  | 2 | GND | Transformer | 053398-0571 | Molex | 1.25mm Pitch SMD CON, Vertical 5-Pin
| 3 | Plasma On Key | Transformer PCB V2.0 | 1 | Plasma On Key
| 4 | GND |  | 2 | GND
| 5 | PULSE_OUT2 |  | 3 | PULSE_OUT2 | Part List
| 6 | PULSE_OUT1 |  | 4 | PULSE_OUT1 | PCB | PN | Vendor | Description
| 7 | +8.0V |  | 5 | +8.0V | Main | DC-011C_SMD |  | DC Power Jack, 0.7 Pie, DC Plug, SMD
[6 empty] | 하네스 | NA645 | 라이트컴 | Coms 컨넥터-DC파워 2.5Ø x 0.7(각형)
| Pin No | Function |  | Pin No | Function |  | 케이블 |  | AWG20
Main PCB V2.0 | 7 | Gas S/V Enable | S/V | 1 | Gas S/V Enable |  | 악어클립 |  | 악어클립 대(검정색)
| 6 | GND |  | 2 | GND
| 5 | Plasma On Key | Transformer PCB V2.0 | 1 | Plasma On Key
| 4 | GND |  | 2 | GND
| 3 | PULSE_OUT2 |  | 3 | PULSE_OUT2
| 2 | PULSE_OUT1 |  | 4 | PULSE_OUT1
| 1 | +8.0V |  | 5 | +8.0V
GND 하네스
Pin-map
| Pin No | Function |  | Pin No | Function
Main PCB | 1 | PAD - GND
| 2 | DC Jack detect Pin
| 3 | Case Pin - GND | GND Cable | 1 | Case Pin - GND
| 4 | Center Pin - GND |  | 2 | Center Pin - GND
| 5 | PAD - GND
| 6 | Hole
| 7 | Hole
JTAG 하네스
Pin-map
| Pin No | Function |  | Pin No | Function
Main PCB | 1 | USB VBUS | ST-LINK | 1,2 | MCU VDD
| 2 | SYS_SWDIO |  | 7 | SYS_SWDIO
| 3 | SYS_SWCLK |  | 9 | SYS_SWCLK
| 4 | USB_ID |  | 3 | GND
| 5 | GND |  | 4 | GND

## ME Size
Battery 용량
| 용량 | W | L | H | 체적
| 2000 | 50 | 60 | 6 | 18000 | SPB605060
| 750 | 48 | 45 | 3 | 6480
|  | 20 | 90 | 10 | 18000
|  | 32 | 51 | 13 | 21216
|  | W | L | H
| PCB | 24 | 68 | 1
| Battery | 20 | 90 | 10
ME | Main | 30 | 100 | 30
| Sub | 30 | 100 | 30
부품 높이 제한
|  | Top | 4 | mm
|  | Bottom | 9 | mm

## Battery CON
|  | Body | Spring | Total
9155-000 | AVX | 3 | 2.3 | 5.3
K113868632 | Top-Link |  |  | 5.3
| Tyco |  |  | 6.2
|  | 3.45 | 1.45 | 4.9
0.089
[5 empty] | 0.204 | 0.182
[5 empty] | 0.115 | 0.11
0.072

## Power review
각 Stage별 전압/전류/파워 | [7 empty] | 각 Block별 소모전류
| Battery | V_batt | 4 | 4 | 4 | V |  |  |  | Current [mA] | Idle |  | Plasma On
|  | I_Peak | 2 | 2.5 | 2.5 | A |  |  |  |  | Status | Current[mA] | Status | Current[mA]
|  | Power_in | 8 | 10 | 10 | W |  | MCU |  | 10 | ON | 10 | ON | 10
| Booster | V_in | 8 | 8 | 8 | V |  | Status LED |  | 7 | ON | 7 | ON | 7
|  | I_in | 1 | 1.25 | 1.25 | A |  | Plasma ON LED |  | 7 | OFF | 0 | ON | 7
| Transformer | Turn ratio | 125 | 392 | 350 |  |  | Power Level LED |  | 18 | ON | 18 | ON | 18
|  | V_out | 1000 | 3136 | 2800 | V |  | Transformer |  | 2500 | OFF | 0 | ON | 2500
|  | I_out | 0.008 | 0.0031888 | 0.0035714 | W |  | S/V |  | 80 | OFF | 0 | ON | 80
[11 empty] | Total | 35 | Total | 2622
Battery 사용 시간 | [7 empty] | 각 mode별 current
| Battery 용량 | 1000 | mAh |  |  |  |  | Idle |  | 35 | mA
| I_Peak | 2 | 2 | 2 | 2 | 2 | A | Plasma On |  | 2622 | mA
| On Time | 20 | 40 | 60 | 80 | 100 | %
| I_average | 400 | 800 | 1200 | 1600 | 2000 | mA | Battery 용량에 따른 사용 시간
| 사용시간 | 2.5 | 1.25 | 0.83333 | 0.625 | 0.5 | Hour | Battery 용량 |  | 1500 | 2000 | 2500 | mAh
[8 empty] | Idle |  | 42.8571 | 57.1429 | 71.4286 | Hour
| Battery 용량 | 2000 | mAh |  |  |  |  | Plasma On | 20 | 171.6247 | 228.833 | 286.0412 | Minute
| I_Peak | 2 | 2 | 2 | 2 | 2 | A |  | 40 | 85.8124 | 114.4165 | 143.0206 | Minute
| On Time | 20 | 40 | 60 | 80 | 100 | % |  | 60 | 57.2082 | 76.2777 | 95.3471 | Minute
| I_average | 400 | 800 | 1200 | 1600 | 2000 | mA |  | 80 | 42.9062 | 57.2082 | 71.5103 | Minute
| 사용시간 | 5 | 2.5 | 1.6667 | 1.25 | 1 | Hour |  | 100 | 34.3249 | 45.7666 | 57.2082 | Minute
| Battery 용량 | 2000 | mAh
| I_Peak | 2.622 | 2.622 | 2.622 | 2.622 | 2.622 | A
| On Time | 20 | 40 | 60 | 80 | 100 | %
| I_average | 524.4 | 1048.8 | 1573.2 | 2097.6 | 2622 | mA
| 사용시간 | 3.8139 | 1.9069 | 1.2713 | 0.95347 | 0.76278 | Hour
| 사용시간 | 228.833 | 114.4165 | 76.2777 | 57.2082 | 45.7666 | Minute

## LED
| Status LED | Plasma LED | Plasma Level LED
| V1.0 | V1.0 | V1.0 | V2.0
VCC | 3.3 | 3.3 | 3.3 | 3.3 | V
R | 430 | 470 | 470 | 180 | Ω
I | 7.6744 | 7.0213 | 7.0213 | 18.3333 | mA

## SV
S/V review data
X605SF
| PN | 912-000007-009
| Test Data - No bias resistor |  |  |  | Press Test @ Vcc=4V
| Vcc[V] | Icc[A] | Operation |  | Press[MPa] | Operation | Icc[A]
| 5 | 0.11 | Open |  | 0.01 | Open | 0.09
| 4.9 | 0.11 | Open |  | 0.02 | Open | 0.09
| 4.8 | 0.11 | Open |  | 0.03 | Open | 0.09
| 4.7 | 0.11 | Open |  | 0.04 | Open | 0.09
| 4.6 | 0.11 | Open |  | 0.05 | Open | 0.09
| 4.5 | 0.1 | Open |  | 0.06 | Open | 0.09 | Plasma 동작범위
| 4.4 | 0.1 | Open |  | 0.07 | Open | 0.09 | Plasma 동작범위
| 4.3 | 0.09 | Open |  | 0.08 | Open | 0.09 | Plasma 동작범위
| 4.2 | 0.09 | Open |  | 0.09 | Open | 0.09
| 4.1 | 0.09 | Open |  | 0.1 | Open | 0.09
| 4 | 0.09 | Open |  | 0.11 | Open | 0.09
| 3.9 | 0.08 | Open |  | 0.12 | Open | 0.09
| 3.8 | 0.08 | Open |  | 0.13 | Open | 0.09
| 3.7 | 0.08 | Open |  | 0.14 | Open | 0.09
| 3.6 | 0.08 | Open |  | 0.15 | Open | 0.09
| 3.5 | 0.08 | Open |  | 0.16 | Open | 0.09
| 3.4 | 0.07 | Open |  | 0.17 | Open | 0.09 | Delay 발생
| 3.3 | 0.07 | Open |  | 0.18 | Closed | 0.09
| 3.2 | 0.07 | Open |  | 0.19 | Closed | 0.09
| 3.1 | 0.07 | Open | Delay 발생 | 0.2 | Closed | 0.09
| 3 | 0.06 | Closed

## Transformer
Transformer Voltage
|  | Tx1 | Tx2 | Tx3 | Unit | Transformer Resistance
Battery | Voltage | 4 | 4 | 4 | V |  | 2차 - 기본 | R_N1 | R_N2 |  |  | 200
| Current | 2.5 | 2.5 | 2.5 | A | 1 | 200 | 0.2 | 363.7 |  |  | 2.41
| Power | 10 | 10 | 10 | W | 2 | 200 | 0.2 | 360.9
Booster | Voltage | 8 | 8 | 8 | V | 3 | 230 | 0.2 | 419
| Current | 1.25 | 1.25 | 1.25 | A | 5 | 230 | 0.2 | 413.2
Transformer | 1차 | 5 | 5 | 5 | N1 | 6 | 230 | 0.2 | 419
| 2차 - 기본 | 200 | 230 | 250 |  | 7 | 250 | 0.2 | 461 | 443 | 장시간 Test 후 dead - 출력이 낮게 나옴
| 2차 total | 1400 | 1610 | 1750 | N2 | 8 | 250 | 0.2 | 458
| Ratio | 280 | 322 | 350 | N1/N2 | 9 | 250 | 0.2 | 460
Output | Voltage | 2240 | 2576 | 2800 | V
| Current | 4.4643 | 3.882 | 3.5714 | mA
Measure | Voltage | 4 | 4
| Current | 2.41 | 2.41
2.33

## Sheet1
AWG Spec
| 저항 |  | 전압 drop | 허용 전류
AWG | Ω/km | Ω/0.1m | 2.5A | 일반 | 실리콘
23 | 66.79 | 0.006679 | 0.016698 | 2.2
24 | 84.22 | 0.008422 | 0.021055 | 0.588 | 5
25 | 106.2 | 0.01062 | 0.02655 | 0.477
26 | 133.9 | 0.01339 | 0.033475 | 0.347 | 3.3
27 | 168.9 | 0.01689 | 0.042225 | 0.288
28 | 212.9 | 0.02129 | 0.053225 | 0.25