## Pipette V1.0
Plasma_Gen_V1.0_20171227.ioc
STM32F070RBT6 |  |  |  |  | TIM Clock
Flash |  | 128 | Kbytes |  | SystemCoreClock | 48000000 | HCLK
SRAM |  | 16 | Kbytes |  | PCLK1 | 48000000 | HCLK / APB1 prescaler
Frequency |  | 48 | MHz | max | TIM clock | 48000000 | PCLK1 x 1
Operating Voltage |  | 2.4 ~ 3.6 | V |  | TIM counter clock | 16000000
Package |  | LQFP64 |  |  | Prescaler | 2 | (TIM clock /TIM counter clock) - 1
[5 empty] | TIM output clock | 80000
[5 empty] | period (ARR) | 199 | (TIM counter clock / TIM output clock) - 1
[5 empty] | Duty | 50
[5 empty] | Pulse | 99.5
[5 empty] | TIM output Clock T | 1.25e-05
PIN No |  | PIN name | Pin Function |  | Cube |  | Pipette V1.0
LQFP64 | I/O |  | Alternate | Additional | Function | GPIO | Off-page | Function | Active | Other | Comment
1 | S | VDD | Digital power supply
2 | TC | PC13 | - | WKUP2 | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
3 | TC | PC14 | - | OSC32_IN | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
4 | TC | PC15 | - | OSC32_OUT | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
5 | FT | PF0 | I2C1_SDA | OSC_IN | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
6 | FT | PF1 | I2C1_SCL | OSC_OUT | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
7 | RST | NRST | Device reset(active low) |  | NRST | - | NRST | Device Reset | LOW
8 | TTa | PC0 | - | ADC_IN10 | GPIO_EXTI0_Rising/Falling | No-PULL | GND_JACK | GND Jack detect | HIGH |  | Low 상태에서 High detect시 GND JACK 인식
9 | TTa | PC1 | - | ADC_IN11 | GPIO_OUT | No-PULL | STA_LED_B | POWER LED-Blue | LOW
10 | TTa | PC2 | SPI2_MISO | ADC_IN12 | GPIO_OUT | No-PULL | STA_LED_G | POWER LED-Green | LOW
11 | TTa | PC3 | SPI2_MOSI | ADC_IN13 | GPIO_OUT | No-PULL | STA_LED_R | POWER LED-Red | LOW
12 | S | VSSA | Analog ground
13 | S | VDDA | Analog power supply
14 | TTa | PA0 | USART2_CTS, USART4_TX | ADC_IN0 | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
15 | TTa | PA1 | USART2_RTS, TIM15_CH1N, USART4_RX, | ADC_IN1 | GPIO_IN | No-PULL | HW_ID0 | HW Version | Open
16 | TTa | PA2 | USART2_TX, TIM15_CH1 | ADC_IN2 | GPIO_IN | No-PULL | HW_ID1 | HW Version | Open
17 | TTa | PA3 | USART2_RX, TIM15_CH2 | ADC_IN3 | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
18 | S | VSS | Ground
19 | S | VDD | Digital power supply
20 | TTa | PA4 | SPI1_NSS, TIM14_CH1, USART2_CK,USB_NOE | ADC_IN4 | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
21 | TTa | PA5 | SPI1_SCK | ADC_IN5 | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
22 | TTa | PA6 | SPI1_MISO, TIM3_CH1, TIM1_BKIN, TIM16_CH1, EVENTOUT, USART3_CTS | ADC_IN6 | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
23 | TTa | PA7 | SPI1_MOSI, TIM3_CH2, TIM14_CH1, TIM1_CH1N, TIM17_CH1, EVENTOUT | ADC_IN7 | ADC_IN7 | No-PULL | BAT_ADC | Battery | -
24 | TTa | PC4 | EVENTOUT, USART3_TX | ADC_IN14 | ADC_IN14 | No-PULL | BATT_THERM | Battery | -
25 | TTa | PC5 | USART3_RX | ADC_IN15, WKUP5 | GPIO_EXTI5_Rising | No-PULL | KEY_PWR | Power On key | HIGH |  | Floating 상태에서 High input시 동작
26 | TTa | PB0 | TIM3_CH3, TIM1_CH2N, EVENTOUT, USART3_CK | ADC_IN8 | GPIO_OUT | No-PULL | PW_CTRL | PS_HOLD | HIGH
27 | TTa | PB1 | TIM3_CH4, USART3_RTS, TIM14_CH1, TIM1_CH3N | ADC_IN9 | GPIO_IN | No-PULL | CHG_STA | Charger | - | ADC_IN9
28 | FT | PB2 | - | -
29 | FT | PB10 | SPI2_SCK, USART3_TX | -
30 | FT | PB11 | USART3_RX, EVENTOUT, I2C2_SDA | -
31 | S | VSS | Ground
32 | S | VDD | Digital power supply
33 | FT | PB12 | TIM1_BKIN, TIM15_BKIN, SPI2_NSS, EVENTOUT, USART3_CK | - | GPIO_OUT | PULL-DOWN | CHG_CTRL | Charger | HIGH |  | Battery Temp over시 High
34 | FTf | PB13 | SPI2_SCK, I2C2_SCL, TIM1_CH1N, USART3_CTS | - | GPIO_OUT | PULL-DOWN | GAS_EN | Gas S/V On Signal | HIGH
35 | FTf | PB14 | SPI2_MISO, I2C2_SDA, TIM1_CH2N, TIM15_CH1, USART3_RTS | - | GPIO_OUT | No-PULL | LED_HIGH | Plsma Level High | LOW
36 | FT | PB15 | SPI2_MOSI, TIM1_CH3N, TIM15_CH1N, TIM15_CH2 | WKUP7, RTC_REFIN | GPIO_OUT | No-PULL | LED_MED | Plsma Level Middle | LOW
37 | FT | PC6 | TIM3_CH1 | - | GPIO_OUT | No-PULL | LED_LOW | Plsma Level Low | LOW
38 | FT | PC7 | TIM3_CH2 | - | GPIO_EXTI7_Rising/Falling | No-PULL | USB_DET | USB | HIGH | GPIO_IN | Floating 상태에서 High input시 동작
39 | FT | PC8 | TIM3_CH3 | - | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
40 | FT | PC9 | TIM3_CH4 | - | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
41 | FT | PA8 | USART1_CK, TIM1_CH1, EVENTOUT, MCO | - | TIM1_CH1 | No-PULL | PULSE_OUT1 | Plasma | HIGH
42 | FT | PA9 | USART1_TX, TIM1_CH2, TIM15_BKIN, I2C1_SCL | - | USART1_TX | n/a | USART1_TX | RS-232 external interface | DATA
43 | FT | PA10 | USART1_RX, TIM1_CH3, TIM17_BKIN, I2C1_SDA | - | USART1_RX | n/a | USART1_RX | RS-232 external interface | DATA
44 | FT | PA11 | USART1_CTS, TIM1_CH4, EVENTOUT | USB_DM | TIM1_CH4 | No-PULL | PULSE_OUT2 | Plasma | HIGH
45 | FT | PA12 | USART1_RTS, TIM1_ETR, EVENTOUT | USB_DP | GPIO_OUT | PULL-DOWN | EN_+8V | Boost | HIGH
46 | FT | PA13 | IR_OUT, SWDIO, USB_NOE | - | SYS_SWDIO | n/a | SYS_SWDIO | SW Download | DATA
47 | S | VSS | Ground
48 | S | VDD | Digital power supply
49 | FT | PA14 | USART2_TX, SWCLK | - | SYS_SWCLK | n/a | SYS_SWCLK | SW Download | CLK
50 | FT | PA15 | SPI1_NSS, USART2_RX, USART4_RTS, EVENTOUT | - | GPIO_IN | No-PULL | USB_ID | USB | HIGH | GPIO_EXTI15_Falling | RS-232를 USB CON으로 사용시 기능 지원 예정
51 | FT | PC10 | USART3_TX, USART4_TX | - | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
52 | FT | PC11 | USART3_RX, USART4_RX | - | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
53 | FT | PC12 | USART3_CK, USART4_CK | - | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
54 | FT | PD2 | TIM3_ETR, USART3_RTS | - | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
55 | FT | PB3 | SPI1_SCK, EVENTOUT | - | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
56 | FT | PB4 | SPI1_MISO, TIM17_BKIN, TIM3_CH1, EVENTOUT | - | GPIO_EXTI4_Falling | PULL-UP | KEY_DOWN | Plasma Level control | LOW |  | Normal High에서 Key input시 Low Active
57 | FT | PB5 | SPI1_MOSI, I2C1_SMBA, TIM16_BKIN, TIM3_CH2 | WKUP6 | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
58 | FTf | PB6 | I2C1_SCL, USART1_TX, TIM16_CH1N | - | GPIO_EXTI6_Falling | PULL-UP | KEY_UP | Plasma Level control | LOW |  | Normal High에서 Key input시 Low Active
59 | FTf | PB7 | I2C1_SDA, USART1_RX, USART4_CTS, TIM17_CH1N | - | GPIO_MODE_ANALOG | GPIO_NOPULL | - | - | -
60 | B | BOOT0 | Boot memory selection |  | GND |  | GND
61 | FTf | PB8 | I2C1_SCL, TIM16_CH1 | - | GPIO_OUT | No-PULL | OP_LED | Plasma On | HIGH | TIM16_CH1 | KEY_OP Input시 상태 변경
62 | FTf | PB9 | SPI2_NSS, I2C1_SDA, IR_OUT, TIM17_CH1, EVENTOUT | - | GPIO_EXTI9_Falling | PULL-UP | KEY_OP | Plasma On | LOW |  | Normal High에서 Key input시 Low Active
63 | S | VSS | Ground
64 | S | VDD | Digital power supply
[5 empty] | GPIO_EXTI0_Rising/Falling | No-PULL | GND_JACK | GND Jack detect | HIGH
[5 empty] | GPIO_EXTI4_Falling | PULL-UP | KEY_DOWN | Plasma Level control | LOW
[5 empty] | GPIO_EXTI5_Rising | No-PULL | KEY_PWR | Power On key | HIGH
[5 empty] | GPIO_EXTI6_Falling | PULL-UP | KEY_UP | Plasma Level control | LOW
[5 empty] | GPIO_EXTI7_Rising/Falling | No-PULL | USB_DET | USB | HIGH
[5 empty] | GPIO_EXTI9_Falling | PULL-UP | KEY_OP | Plasma On | LOW

## TIMER
System clock set-up
Plasma_Gen_V1.0_20171227.ioc
System 관련 clock set-up 부분으로 STM32Cube에서 설정한다.
code 위치 : main.c의 void SystemClock_Config(void)에서 설정함
1) PREDIV, PLLMUL, AHB Prescaler를 통해 SystemCoreClock을 설정한다.
|  |  | STM32F072RB-Nucleo | new | test | old | TIM3 | Unit | Comment
|  | HIS | 8000000 | 8000000 | 8000000 | 8000000 | 8000000 | Hz | Internal RC Frequency
|  | PREDIV | 2 | 2 | 2 | 1 | 2
|  | VCO input | 4000000 | 4000000 | 4000000 | 8000000 | 4000000 | Hz | PLL VCO input Frequency
|  | PLLMUL | 12 | 12 | 12 | 6 | 12
|  | SYSCLK | 48000000 | 48000000 | 48000000 | 48000000 | 48000000 | Hz | System Core Clock
|  | AHB Prescaler | 1 | 1 | 1 | 1 | 1
|  | HCLK | 48000000 | 48000000 | 48000000 | 48000000 | 48000000 | Hz
2) HCLK를 APB1 Prescaler로 나눠서 PCLK1을 계산하고, APB1 Prescaler value에 따라 자동으로 설정되는 APB1 Timer scaler에 의해 APB1 Timer clock이 계산된다.
APB1 Timer clock이 Timer의 input clock이 된다.
|  | APB1 Prescaler | 1 | 1 | 1 | 4 | 1
|  | PCLK1 | 48000000 | 48000000 | 48000000 | 12000000 | 48000000 | Hz | APB1 peripheral clock
|  | APB1 Timer scale | 1 | 1 | 1 | 2 | 1 |  | APB1 Prescaler에 의해 자동 설정되는 항목으로 cube로 source 생성시 변경되는 부분이 없는 항목임
|  | APB1 Timer clock | 48000000 | 48000000 | 48000000 | 24000000 | 48000000 | Hz | PCLK1 x APB1 Timer scale(TIM input clock) | TIM2 input clock (TIM2CLK) is set to APB1 clock (PCLK1), since APB1 prescaler is equal to 1
TIMER Clock Set-up
code 위치 : tim.c의 void MX_TIM1_Init(void)에서 설정함
3) Timer의 counter clock을 설정하는 부분으로 TIM counter clock를 결정하고 이에 맞는 Prescaler 값을 찾아 설정한다.
|  | TIM counter clock | 16000000 | 16000000 | 8000000 | 8000000 | 1000000 | Hz | target value
|  | Prescaler | 2 | 2 | 5 | 2 | 47 |  | (APB1 Timer clock / TIM2 counter clock) - 1
4) Timer의 ouput clock을 설정하는 부분으로 위의 TIM output clock를 결정하고 이값으로 TIM counter clock을 나눠서 그에 맞는 Period 값을 계산하여 설정한다.
|  | TIM output clock | 24000 | 80000 | 80000 | 80000 | 80000 | Hz
|  | Period(ARR) | 665 | 199 | 99 | 99 | 11 |  | (TIM counter clock / TIM2 output clock) - 1
|  | TIM output clock_min | 244.1406 | 244.1406 | 122.0703 | 122.0703 | 15.2588 | Hz | Period가 Max 65535일때 TIM output clock frequency
|  | TIM output clock_max | 16000000 | 16000000 | 8000000 | 8000000 | 1000000 | Hz | Period가 Min 0일때 TIM output clock frequency -> TIM counter clock과 동일함
|  | Duty cycle | 50 | 50 | 50 | 50 | 50 | %
|  | Pulse | 332 | 99 | 49 | 49 | 5 |  | Period(ARR) x duty / 100
0.0125
0.000125
100