## Issue list
No | Issue Date | Status | Closed Date | Issue | Command
1 | 2018-05-26 | Closed | 2018-05-26 | 초기에 모든 LED가 ON 상태가 된다. - STA_LED_B_Pin¦STA_LED_G_Pin¦STA_LED_R_Pin¦LED_LOW_Pin - LED_HIGH_Pin¦LED_MED_Pin¦OP_LED_Pin | gpio.c 수정 HAL_GPIO_WritePin(GPIOC, LED_LOW_Pin, GPIO_PIN_SET); HAL_GPIO_WritePin(GPIOB, LED_HIGH_Pin¦LED_MED_Pin¦OP_LED_Pin, GPIO_PIN_SET);
2 | 2018-05-26 | Pending |  | Low Battery 상태에서 Plasma disable 시킴 - 현재 Low battery에서 Plasma On 됨 | 적용하지 않기로 함 - 추후 필요시 적용 여부 판단
3 | 2018-05-26 | Closed | 2018-06-07 | 충전중 Plasma disable 시킴 | gpio.c 수정 HAL_GPIO_EXTI_Callback()의 Key interrupt에 아래 조건 추가 if( HAL_GPIO_ReadPin(USB_DET_GPIO_Port,USB_DET_Pin) == GPIO_PIN_RESET)
4 | 2018-05-26 | Closed | 2018-06-07 | Plasma On 상태에서는 Level control disabel 시킴 | gpio.c 수정 HAL_GPIO_EXTI_Callback()의 Key interrupt에 아래 조건 추가 if(HAL_GPIO_ReadPin(KEY_OP_GPIO_Port,KEY_OP_Pin) == GPIO_PIN_SET)
5 | 2018-05-26 | Closed | 2018-06-07 | booting 초기에 Plasma Level를 Low로 Setting하고 LOW_LED On 상태를 유지함 - Plasma 동작전에 Level를 Setting함. | Power On LED Cycling 후 동작 시작하도록 수정 Chargetask_init.c OutEventtask_signals.h mc.c
6 | 2018-05-26 | Closed | 2018-06-04 | PWM Pulse frequency 수정 80KHz -> 70KHz | Plasma_state.c 수정 #define TIM1_OUTPUT_CLK (70000) //6/4 Arvid TIM1 output clock = 70KHz
7 | 2018-05-26 | Closed | 2018-06-05 | Plasma Level control 안됨 - Always On으로 동작함 - Plasma On time 계산 수정 : 4msec / 6msec / 8msec | Plasma_state.c 수정 #define PLASMA_STEP_TIME_PWM_MSEC 2 PlasmaOnTime = PLASMA_STEP_TIME_PWM_MSEC*(PlasmaLevel+2); uint32_t getPlasmaVolLevel() 삭제
8 | 2018-05-26 | Closed | 2018-06-07 | Plasma On 상태에서만 Level control Key가 동작함 | InEventtask.c - DOWN/UP Key event에 아래 code 추가 xTimerStop(xkeyCheckTimer,0); xTimerStart(xkeyCheckTimer,0);
9 | 2018-06-07 | Closed | 2018-06-07 | PWM Pulse frequency 낮게 나옴 - 30KHz 정도 | main.c 수정 RCC_ClkInitStruct.APB1CLKDivider = RCC_HCLK_DIV1;
10 | 2019-01-01 | OPEN |  | 완충 인식 안됨 | InEventProcess.c else if(adc_val >= CHARGEADC_HIGH) //charge complete { if(IsExtPwrInState == TRUE) { if(IsCharging != 1) // 2019.01.22 Arvid Full Charging detect error
11

## Function
Fucntion |  | 하부 Item | 동작 | P/F | Commant
Power | Battery | normal Status | LED Blue On | OK | 3.6V 이상
|  | Low battery Status | LED Red blinking On | OK | 3.5V
|  | Charging | LED Red On | OK
|  | Full Charge | LED Green On | -
Plasma Level | Initial status | Level | Level = Low | OK | Booting 초기에 무조건 Low level로 Setting 되어야 함
|  | LED | LED_LOW On | OK
| UP Key | Level | 현재 상태에서 상위 Level로 Set | OK | High Level에서 UP Key 입력이 들어와도 무시함
|  | LED | 현재 상태에서 상위 Level의 LED On | OK
| DOWN Key | Level | 현재 상태에서 하위 Level로 Set | OK | Low Level에서 DOWN Key 입력이 들어와도 무시함
|  | LED | 현재 상태에서 하위 Level의 LED On | OK
| 예외처리 | Plasma On 상태에서는 Level control disabel 시킴 |  | OK
Plasma | Normal Status | Key input | Low detect - Push | OK
|  | LED | Push중 On 상태 유지 | OK
|  | Pulse | Push중 On 상태 유지 | OK
|  | S/V | Push중 On 상태 유지 | OK
| On/Off time | Low level | On : 4msec / Off : 6msec | OK
|  | Middle Level | On : 6msec / Off : 4msec | OK
|  | High Level | On : 8msec / Off : 2msec | OK
| Low Battery | Plasma Function disable
| Charging Status | Plasma Function disable |  | OK

## Sheet3
On/Off time | Low level | On : 4msec / Off : 6msec | PLASMA_SPAN_TIME_PWM_MSEC | 10 | msec
| Middle Level | On : 6msec / Off : 4msec | PLASMA_STEP_TIME_PWM_MSEC | 2 | msec
| High Level | On : 8msec / Off : 2msec
|  |  | Level | On Time
|  |  | 0 | 4
|  |  | 1 | 6
|  |  | 2 | 8

## Sheet1
mc.c
mctask_init()
mcCLIregister();
check_pwronstate();
sendTaskStartSignal();
sendtoBuzzerOnCmd(1);
Chargetask_state.c
CheckIsChargeCmd(command_type *command)
status_led_ctr(boolean onoff, uint8_t chan)
switch(command->cmd)
{
case CHARGE_TASK_STOP_F:
charge_state = CHARGESTATE_TEST;
charge_state_machine(command);
break;
case CHARGE_STATUS_ON_F:
status_led_ctr(ON, (uint8_t)data[FDPOS(charge_cmd_msg_field_type,Cmd)]);
DBGHI(OUTEVENT, "Status LED (%d) ON.\n", (uint8_t)data[FDPOS(charge_cmd_msg_field_type,Cmd)] );
break;
case CHARGE_STATUS_OFF_F:
status_led_ctr(OFF, (uint8_t)data[FDPOS(charge_cmd_msg_field_type,Cmd)]);
DBGHI(OUTEVENT, "Status LED (%d) OFF.\n", (uint8_t)data[FDPOS(charge_cmd_msg_field_type,Cmd)] );
break;
case CHARGE_STATUS_ALL_OFF_F:
STATUS_LED_OFF(STA_LED_R);
STATUS_LED_OFF(STA_LED_G);
STATUS_LED_OFF(STA_LED_B);
DBGHI(OUTEVENT, "All Status LED OFF.\n" );
break;
default:
charge_state_machine(command);
break;
}
charge_state_machine (command_type *cmdptr)
static void vMCTask( void *pvParameters );
static void prvmcTimerCallback( TimerHandle_t pxExpiredTimer )
void vStartMCTasks( void )
case 0:
// No cmd rxed, 2000 msec wait timeout
case MC_POWEROFF_CMD_F:
sendSTOPandPowerOFF();
case MC_PWR_KEY_CHECK_TIMER_F:
sendMcCommand(MC_POWEROFF_CMD_F);
case PWR_KEY_PRESSED_EVT:
StartCheckPwrKey();
case MC_LOWBATT_OCCURR_EVT:
sendtoOutEventCmd(CHARGE_LOWBATT_OCCURR_EVT);
case MC_CHARGING_STARTED_EVT:
sendtoOutEventCmd(CHARGE_CHARGING_STARTED_EVT);
case MC_FULL_CHARGE_EVT:
sendtoOutEventCmd(CHARGE_FULL_CHARGE_EVT);
case MC_CHARGING_STOPED_EVT:
sendtoOutEventCmd(CHARGE_CHARGING_STOPED_EVT);
case MC_EXTPWR_IN_EVT:
sendtoOutEventCmd(CHARGE_EXTPWR_IN_EVT);
case MC_EXTPWR_OUT_EVT:
sendtoOutEventCmd(CHARGE_EXTPWR_OUT_EVT);
case MC_OPERATE_ON_F:
sendtoOutEventCmd(PLASMA_LED_ON_F);
sendtoOutEventCmd(PLASMA_START_CMD_F);
case MC_OPERATE_OFF_F:
sendtoOutEventCmd(PLASMA_LED_OFF_F);
sendtoOutEventCmd(PLASMA_STOP_CMD_F);
case MC_DOWN_KEY_PRESSED_F:
case MC_UP_KEY_PRESSED_F:
PlasmaVolSet(cmd.cmd);
case MC_GND_JACK_IN_EVT:
case MC_GND_JACK_OUT_EVT:
default:
break;
mc_process.c
void sendTaskStartSignal()
void check_pwronstate()
void StartCheckPwrKey(void)
boolean check_pwroffstate(void)
void sendSTOPandPowerOFF(void)
void PlasmaVolSet(uint16_t command)